# Technical Report

Magnetic Gadget Kit

Designed and Produced by Haixiang Xu @ Meister Lab

Sep. 2016

## Overview

The project is intended to design and produce a set of lab experiment kit for magnetic perception research of neurons. It includes two major parts: the electromagnet and the magnetometer. The design files and user manuals are also included.

The electromagnet, also referred to as ‘Magneto’, is a 3.5mm audio jack powered, compact gadget to generate static or low frequency alternative electromagnetic fields, which is the most important part of generating magnetic field necessary for researches on magnetic field perception. And the magnetometer provides the additional, accurate measurement to properly monitor and record the field that is actually applied to the neurons.

Figure 1. Magneto gadgets				Figure 2. Magnetometer

There are 50 electromagnets and 1 magnetometer already made and ready to be implemented, the design files and production guidance is written up for future manufacturing. Please read through this report thoroughly before using these gadgets or making any modification of these products. The whole project files can be found on the following link:

https://github.com/lx9t01/MagnetGadget.git

## Electromagnet

- Design and Production

The design of Magneto is straightforward. It is bobbin shaped, and consists of two coaxial plastic cylinders which act as inner tube and outer case. On each end of the cylinder there are epoxy seals that help fixing the bobbin and the coil inside.

The cylinders are produced with 3D printing, and designed with Autodesk Inventor. The models are tuned with Cura slicing software, and the model files can be found in the package. 3D printers in Caltech TechLab is free to use, and we recommend using Bukito or Mod-T to print these materials.

Inside the inner cylinder, we insert an iron core to concentrate and enhance the magnetic field.

Fig 3. The three-view of Magneto

The design drawing can be found in the ‘./3D Printing Design’ folder, and the plastic casing that is being printed is shown below:

Fig 4. The design drawing, 3D effect view

Fig 5. The printed casing

In between the coaxial cylinders there is the magnetic coil. The coils are made from 30 AWG magnetic wires, and in each coil there are 1000 windings of such wires.

Fig 6. Break-out look inside the bobbin

The ends of the coil are soldered to the break-outs of a 3.5mm audio cable. Note that only the left audio signal channel and the ground line are used, the right channel is open circuited. The junction is fixed and strengthened by head-shrinking tubes to ensure proper mechanical strength.

- Properties

Physical Specs:

| Inner cylinder | Inner diameter | 8 mm |
| --- | --- | --- |
|  | Outer diameter | 11 mm |
| Outer cylinder | Inner diameter | 21 mm |
|  | Outer diameter | 24 mm |
| End wall thickness |  | 2 mm |
| Total height |  | 24 mm |
| Working temperature |  | 0 ~ 50 ºC |
| Audio cable length |  | 50 ~ 60 cm |

Electrical Properties:

| Resistance | 17 ± 1 Ω |
| --- | --- |
| Coil winding | 1000 round |
| Input max voltage | 2V peak to peak sine wave |
| Input max current | 120 mA |
| Max frequency | 100 Hz |
| Recommended working distance | 1 ~ 4 cm along the central axis |

- Usage

The usage of the gadget is extremely easy. The gadget is compatible with most smartphones, both in Android and iOS systems.

  - Plug the audio cable into 3.5mm audio jack

  - For alternative magnetic field, use signal generating Apps to generate waveform to the audio output

    - Signal Generator: Audio Test Tone Utility

https://itunes.apple.com/us/app/signal-generator-audio-test/id543661843?mt=8

    - Function Generator

https://play.google.com/store/apps/details?id=com.keuwl.functiongenerator&hl=en

  - For other special waveform or static magnetic field, please use proper app or use a professional function generator

  - Tune the smartphone to proper volume, and the magnetic field is ready to be used

  - Plug, clipper or attach the gadget near the neuron or object that you intend to work on.

- Manufacturing steps

The manufacturing or production of the gadget depends on many software, the general rules follow ‘design formatsliceprint’ pattern. To make more gadgets in the future, please follow the following steps:

  - Download and install Autodesk Inventor, or Windows 10 pre-installed 3D Builder software

  - Open design files (.dmg), modify them and save as .stl file

  - Open the .stl file with Cura, properly set Cura with specific printer parameters, and place the model in workspace

  - Modify in Cura, print in batches or place models closer, save the file with .gcode format or .g format

  - If use Mod-T, connect the machine with USB cable,

    - Open .gcode file with Mod-T printer utility

    - Load .gcode files, press print to start printing

  - If use Bukito,

    - Change name of the .gcode file to auto0.g

    - Copy the file into micro SD card, insert into Bukito machine, press start to begin printing

  - After the 3D bobbin is printed, wind the magnetic wire around the inner cylinder with the winding machine for 1000 rounds

  - After the coil is finished, insert the coin into the outer case, leaving the ends of magnetic wire out from the end seals

  - Peel the audio cable and get the left channel, ground line

  - Solder the ends of wire to the break-outs of audio cable, before soldering you might need to polish the end of the wires so as to get rid of the isolation paint from the wire

  - Using heat-shrinking tube to wrap the connection for isolation and strengthen purposes

Fig 7. Manual winding machine

For the detailed training of 3D printing process or design procedures, please follow the instructions in the link below, or ask for assistance from TechLab librarians.

How To: use Cura with Mod-T

http://blog.newmatter.com/how-to-use-cura-with-the-mod-t/

Bukubot website

http://bukobot.com/bukito

## Magnetometer

- Design and Overview

The magnetometer is an Arduino based magnetic field sensor. We choose MAG3110 as magnetic sensor chip, and use Arduino Fio v3 as data board. The system is powered with USB cable and transmit data back to PC with USB-Serial port. Data transmission between data board and sensor is via I2C protocol. To make the magnetometer more compatible and easy to use, we designed a 3D printed case to protect the data board, connecting MAG3110 sensor with Dupont wires.

Fig 8. Magnetometer overview

Fig 9. Magnetometer case design

All the sensor chips and Arduino board can be purchased from sparkfun website, and made easily with the software provided in the package. A Python script is used for PC or Mac to collect data from USB port, data will be stored in a plain text file, which can be imported with Matlab or other data analysis tools.

The magnetometer is equipped with auto-calibration features, the first 15s after powering up is used explicitly for calibration purpose. To properly calibrate the sensor, please put the sensor inside the Zero-field cylinder while the blue LED light is blinking. After the blue light stabilized, take out the sensor and it is ready to measure magnetic field. The magnetometer can be used for a wide range of magnetic fields. Data collection rate is set to 50 Hz, however, this rate can be further tuned according to the later using scenarios.

Moreover, the script is featured with a graph display of real-time magnetic field amplitude for more convenient magnetic field monitoring.

Fig 10. Graphical UI monitor

And the data is recorded as:

Time (ms)		x		y		z	 total

21500 16.20 34.80 -17.70 42.27

21520 15.90 34.10 -17.70 41.58

21540 15.80 34.70 -17.60 41.99

21560 16.40 34.60 -16.50 41.69

21580 16.20 34.90 -18.50 42.69

- Hardware List

| Ardiono Fio v3 | 1 |
| --- | --- |
| MAG3110 breakout chip | 1 |
| USB cable | 1 |
| Dupont wire (F-F) | 4 |
| 3D printed case | 1 |
| Zero-field calibration cylinder | 1 |

- Software/Script List

| Arduino design file | Mag3110_Fiov3.ino |
| --- | --- |
| Python data script | USB_Serial_Data_Acquisition.py |

- Specs and Features

| Power supply | 5V (USB) |
| --- | --- |
| Vcc | 3.3V (Fio v3) |
| Magnetic field range | 1000 mT |
| Number of measuring axis | 3 |
| Communication Protocol | I2C |
| Data collection rate | 50Hz |
| Accuracy | 0.1 mT |
| Calibration period | First 15s powering up |
| Length | 69 mm |
| Width | 32 mm |
| Height | 22 mm |

- Usage

  - Keep the sensor in the Cylinder

  - Plug in USB cable to power up

  - Wait for the magnetometer to calibrate, blue LED will blink. During calibration, run python script with terminal call

python ./USB_Serial_Data_Acquisition.py

  - After the LED lights up, take out the sensor and measure. Monitor the field amplitude from graphic UI

  - After the measurement, put the sensor back to the zero-field cylinder

  - Unplug the USB cable, or press Ctrl-C to end the Python script

  - Save result from result.txt and use for future analytics

- Production

The production of magnetometer is very simple. Purchase any of the compatible Arduino boards online, and at the first time of use, plug in the USB cable, download .ino program from Arduino Software to your board. After the download is complete, Arduino board will run the program automatically.

Copy the python file in the same folder, and it can be used on any PC machine with Python environment installed. Run the python script while the board is calibrating, and it’s good to measure.

Haixiang Xu: hxu@caltech.edu

http://www.autodesk.com/products/inventor/overview

https://www.microsoft.com/en-us/store/p/3d-builder/9wzdncrfj3t6

https://ultimaker.com/en/products/cura-software

https://store.newmatter.com/#!/setup4

https://www.sparkfun.com/

According to MAG3110 datasheet     https://cdn.sparkfun.com/datasheets/Sensors/Magneto/MAG3110_v9.2.pdf

https://www.arduino.cc/en/Main/Software